Instructions for completing the Haigis constant optimization Excel spreadsheet for optical biometry:
The Excel spreadsheet "Haigis-300.xls" is for collecting the data necessary to optimize the a0, a1 and a2 constants for the Haigis formula using optical biometry. Please provide data for a minimum of 200 patients (250 will give the best outcome) who have all had the same intraocular lens implant. The axial length range for your data series should be as broad as possible, with a range from at least 20 mm to 28 mm being ideal. 

Data

The Patient Identifier should be a HIPPA compliant reference that will allow you to identify this particular case in the event that a data entry must be double checked. Axial Length is entered in mm. Phakic ACD is the measurement in mm from the corneal vertex to the anterior lens capsule. If you are using the Haag-Streit Lenstar, the Phakic ACD is the same and the central corneal thickness + the aqueous depth. Do not simply enter the Lenstar “AD” (aqueous depth) or the optimized lens constants produced will give consistent hyperopia. Pre-op K1 and K2 are the power in diopters of the two principal meridians obtained by keratometry. IOL Power is the power of the intraocular lens implanted at the time of surgery in diopters. Target SE does not need to be entered. This is not part of the calculation process and is offered simply as a way of tracking what the intended refractive objective was for each individual case. Post-op Sphere and Post-op Cyl are the components of the final, stable post-operative manifest refraction. Auto-refractions should not be entered into these fields. 
Which cases should I include?

For these patient entries, enter cases where you are certain of the quality of the pre-operative measurements, uncomplicated surgery and those for which a clear, stable endpoint manifest refraction was possible. In other words, cases that reflect your best work. 
You should exclude cases were the pre-operative measurements were in question, those with complications during surgery, such as: capsular tears, sulcus fixated lenses, aphakic and pseudophakic IOL exchanges, indwelling silicone oil, prior retinal detachment, etc. and those for which the final refraction is less than optimal. For refractive stability, it is important to include cases only where the capsulorhexis is round, smaller than the optic, and centered. 
Best-corrected visual acuity should be 20/40, or better and the refraction should be stable. For the sake of this exercise, a three to ten week postoperative refraction is considered to be stable. It is very important that cases only be excluded for pre-operative and intraoperative issues and not for a refractive outcome that is different than anticipated. A data set that has been edited for refractive outcomes may significantly skew the optimized lens constants. 

For optimization of the a0, a1 and a2 Haigis constants, the blue fields must be filled in and the tan fields left alone, as they contain data, or internal formulas and/or Excel macros. It is not necessary to fill in the final post-operative visual acuity, or the target refraction (green fields) for this exercise.

Do not make changes in the spreadsheet

Please do not enter anything except data into the blue fields, and please be careful not to change anything on the spreadsheet itself. Deleting rows, adding columns, or making format changes may disrupt the internal mathematics of the spreadsheet running invisibly in the background. It is also important to double-check all entries, as a single entry error (such as transposed digits) may corrupt the final result. In our office, we usually have two people review all data entries for accuracy.
What do I do with the spreadsheet?
When you have entered as many cases as possible, please send it back to Dr. Hill at hill@doctor-hill.com as an e-mail attachment. Dr. Hill will then carefully run a 3-variable regression analysis, calculating the a0, a1 and a2 constants for this particular IOL. Some routine statistical filtering of the data may exclude a handful of cases that fall well outside an initial prediction model. Using this same data, lens constants will also be optimized for the Holladay 1, SRK/T and Hoffer Q formulas. 
If, following data filtration, there remain less than 200 cases, only the Haigis a0 constant will be optimized. In order for all three Haigis constants to be fully optimized, Dr. Haigis requests that a minimum of 200 valid cases must be available for analysis. For this reason, submitting 250 cases is ideal. 

If submitting IOLMaster data for analysis, it is assumed that we have your permission to post the outcomes to Users Group for Laser Interference Biometry (ULIB) web site for this particular IOL.

              http://www.augenklinik.uni-wuerzburg.de/eulib/index.htm
If submitting Lenstar data for analysis, your cases for a particular IOL model will be pooled with those from other surgeons to generate a large data set from which Lenstar constants will be derived. 

Please ignore any statistical information at the bottom of the spreadsheet and any divide by zero statements. These areas will be updated after your cases have been reviewed. 
There is space available on the spreadsheet for 300 cases. It is not required that you enter that many cases, but try to provide as many as possible. The more cases, and the wider the axial length range, the better the overall accuracy.

If you have any questions, feel free to contact Dr. Hill at hill@doctor-hill.com
June, 2010
